Rotary Evaporator EV400

Technical Specifications:

- Lifting Mechanism: Motorized automatic lift; lifting stroke: 160 mm

- ★ Rotation Speed Range: 20–300 rpm; LCD digital display for clear and stable readout

- Condenser: Vertical serpentine condenser, 1200 cm² standard condensation area; optional 1500 cm² condenser available

- Adjustable angle of rotary joint: 0–60°

- Automatic lift upon power failure to separate evaporating flask from heating bath, preventing localized overheating

- ★ Secondary condensation recovery system ensures high sample recovery and protects vacuum pump from solvent corrosion

- Evaporating Flask Capacity: 50 mL – 3000 mL

- ★ Heating Bath Temperature Range: Room temperature to 210 °C, switchable between water/oil bath

  - Bath capacity: 5 L

  - Bath material: 304 stainless steel; optional PTFE coating

- Temperature Control: Intelligent PID control (LCD display)

  - Temperature accuracy: ±1 °C (water), ±2 °C (oil)

- Dry-boil protection with automatic power-off

- Over-temperature alarm for water bath

- Timer function for heating bath

- Optional digital vacuum controller for precise vacuum regulation during distillation

Standard Configuration:

- Rotary evaporator main unit ×1

- Heating oil bath ×1

- Vertical serpentine condenser ×1

- Gas-liquid separation flask ×1

- Evaporating flask ×1 (1000 mL)

- Receiving flask ×1 (1000 mL)

Click the card above to download the translated Word document. Let me know if you'd like both EV400H-L and EV400-V versions combined or formatted for a product catalog.