# Rotary Evaporator EV400 Touch-L

Technical Specifications:

▲ Lifting mechanism: Motorized automatic lift, lifting stroke 160 mm
▲ Rotation speed range: 20–300 rpm
Condenser: Vertical coil condenser, condensation area 1200 cm²
Adjustable rotation head angle: 0–60°
Automatic lift when power is off, separating the evaporating flask from the bath to prevent localized overheating of the sample
▲ Equipped with a secondary condensation recovery system to ensure sample recovery rate and protect the vacuum pump from solvent corrosion
Evaporating flask capacity: 50 mL–3000 mL
Bath temperature range: Ambient to 210 °C, switchable between water and oil bath; bath capacity: 5 L; bath material: 304 stainless steel
Bath temperature control: Intelligent PID control (LCD display); temperature accuracy: ±1 °C (water), ±2 °C (oil)
Bath equipped with dry-heating protection and automatic power-off function; water bath high-temperature audible alarm
▲ 4.3-inch color touchscreen controller displaying multiple parameters such as rotation speed simultaneously, with built-in graphical distillation curves
▲ Gradient distillation mode with programmable automatic speed control
Intermittent forward and reverse rotation function
USB data interface for easy export of experimental data
▲ Optional digital vapor temperature display for precise monitoring of sample vapor temperature
Intelligent timed distillation function

Configuration:

- Rotary evaporator main unit ×1 (with color touchscreen controller ×1)
- Heating oil bath ×1
- Inclined coil condenser ×1
- Gas-liquid separation flask ×1
- Evaporating flask (1000 mL) ×1
- Receiving flask (1000 mL) ×1