Fully Automatic Digestion Unit AutoDigiBlock S30 UP Parameters

1. Purpose:

Used for the fully automatic wet digestion of various samples (fully automatic reagent addition, automatic mixing, automatic lifting/lowering, automatic programmed temperature ramp digestion, automatic acid driving, automatic volume adjustment, etc.).

2. Technical Parameters:

2.1 Heating Module

- Temperature Control Range: Room temperature ~ 230°C

- ★ Features pre-heating function, effectively shortening experiment time and improving efficiency.

- Control Method: Intelligent PID control, temperature controller is originally imported; Control Accuracy: ±0.1°C

- ★ Sample Capacity: 30 positions × 50mL; Digestion hole depth between 4cm and 6cm, ensuring heating effectiveness while efficiently achieving reagent condensation reflux, reducing both reagent consumption and background values.

- Surround-style graphite heating body, surface treated with Teflon coating.

- Digestion process allows for over 20 levels of programmed temperature ramps; Settable heating rate facilitates digestion control.

- Users can remotely monitor digestion temperature, time, system status, and progress using a mobile phone or office computer.

2.2 Fully Automatic Reagent Addition System

- Reagent Channels: 6, supporting automatic switching between 6 digestion reagents.

- Reagent addition speed: 2mL/s, adjustable from 0-5mL/s; Liquid addition accuracy: better than 1%.

- Reagent addition uses high-precision peristaltic pumps; Reagent delivery lines are PFA material; All channels can safely handle various corrosive reagents (including Hydrofluoric Acid).

- ★ XY-axis全方位移动机械臂, mechanical arm transmission lead screw, precise positioning, fully enclosed robotic arm/transmission parts have zero contact with acid vapors/liquids, beltless consumable design, ensuring long service life.

- Tilted liquid addition tube design, reagents flow down the digestion tube wall, effectively reducing powder sample adhesion and preventing reagent/sample splashing.

2.3 Fully Automatic Lifting & Oscillation System

- ★ Fully automatic oscillation mixing function, using a motor oscillation system and digestion tube rack to mix reagents and samples; entire process is contact-free, preventing contamination and cross-contamination; Can perform mechanical oscillation mixing on 30 samples simultaneously.

- Oscillation power continuously adjustable, oscillation time settable; Multiple mixing modes available.

- Automatic lifting function raises the tube rack before acid addition and after acid driving, ensuring tube bottoms completely clear the heating module; Settable cooling time during lifted acid addition ensures safety.

2.4 Automatic Sample Volume Adjustment System

After digestion and acid driving, volume adjustment is performed at room temperature using short-range, low-power ultrasonic sensors to measure volume in the tubes; High-precision peristaltic pumps adjust samples to the specified volume (10-50mL); Adjustment accuracy: better than 1%.

2.5 Automatic Acid Fume Exhaust System

- Integrated automatic acid fume exhaust system, can be used without needing a fume hood. Contains fumes within the system, preventing damage to electronic components when placed entirely inside a hood.

- ★ Exhaust airflow can be automatically adjusted based on experimental steps and sample volume.

- ★ Standard exhaust system includes HEPA-grade filtration, continuously purifying system air during digestion to avoid sample contamination.

- Exhaust system includes automatic condensate drainage.

- ★ Compact instrument design, exhaust system is detachable, allowing the unit to be used inside a fume hood after detachment.

2.6 Intelligent Control Software

- PC-controlled, powerful functionality; Software supports easy switching between Chinese/English; Main interface shows reagent levels for 6 channels with alarms, and waste bottle full warning, with user-definable limits.

- ★ Scheduled start function: Methods can be set to run at a predefined time, the instrument starts digestion automatically, improving efficiency.

- ★ Voice prompt function: Notifies user upon method completion, preventing samples from drying due to residual heat, enhancing safety and reliability.

- ★ Wireless computer connection, remote control capable, computer can be placed outside the laboratory ventilation system to control the main unit.

- ★ Remote monitoring system: In a network environment, users can monitor real-time video of the experiment process from outside the lab (e.g., office) via remote terminals, tracking progress remotely; Allows remote emergency stop.

- Freely set and store digestion and other pretreatment steps to meet various project requirements.

- Automatically generates digestion process reports.

3. Configuration List:

- One fully automatic graphite digestion main unit;

- One set of automatic reagent addition system;

- One set of anti-corrosion graphite heating system;

- One set of oscillation mixing system;

- One set of automatic acid fume exhaust system with HEPA purification;

- Digestion tube racks (30-position × 1);

- One set of HEPA-grade filter purification mesh;

- One set of PTFE digestion tubes (30 tubes);

- One set of PP polypropylene digestion tubes (500 tubes);

- One set of intelligent control software;

- One brand laptop (Win 10 OS, 4G RAM, 500G HDD, 14");

- One set of remote monitoring system software; Optional

4. Technical Service & Instrument Training

- ★ The bid product must be mature, with over 400 domestic users; must provide over 10 contracts and a list of 400+ users.

- Engineers will install and debug the instrument at the user's lab upon notification until acceptance is passed.

- On-site free training for 2-5 technical personnel until independent operation is achieved, covering principles, debugging, operation, and maintenance.

- Manufacturer maintains an application lab domestically, providing free lifelong application technical consultation.

- ★ Manufacturer is ISO9001 certified, has an after-sales service center in ______, and responds within 24 hours to user repair requests.