Fully Automatic Digestion Unit D-DOCTOR Parameters

1. Purpose:

Used for the fully automatic wet digestion of various samples (fully automatic reagent addition, automatic mixing, automatic lifting/lowering, automatic programmed temperature ramp digestion, automatic acid driving, automatic volume adjustment, etc.).

2. Technical Parameters:

2.1 Heating Module

- Temperature Control Range: Room temperature ~ 230°C

- ★ Features pre-heating function, effectively shortening experiment time and improving efficiency.

- Control Method: Intelligent PID control; Control Accuracy: ±0.1°C

- ★ Sample Capacity: 84 positions × 50mL; Digestion hole depth between 4cm and 6cm, ensuring heating effectiveness while efficiently achieving reagent condensation reflux, reducing both reagent consumption and background values.

- ★ Precise sample observation area, equipped with base lighting, allows for accurate observation of sample digestion status and remaining reagent levels.

- Surround-style graphite heating body, surface treated with Teflon coating; includes two independent 42-position graphite heating blocks, allowing different temperatures for digesting different samples.

- Digestion process allows for over 20 levels of programmed temperature ramps, facilitating digestion control.

- Users can monitor digestion temperature, time, system status, and progress via the laboratory LAN using a mobile phone or office computer.

2.2 Fully Automatic Reagent Addition System

- ★ Reagent Channels: 10, supporting automatic switching between 10 digestion reagents; optionally equipped with high-corrosion-resistant rotary valves to expand channels up to 20.

- ★ Peristaltic pump reagent addition speed: 2mL/s, adjustable from 0-5ml/s; Liquid addition accuracy: better than 1%; Syringe pump liquid addition accuracy: better than 0.1%.

- ★ Reagent addition utilizes one high-precision peristaltic pump and one high-precision syringe pump; Reagent delivery lines are PFA material; All channels can safely handle various corrosive reagents (including Hydrofluoric Acid).

- ★ Uses a 360° rotary robotic arm with curved liquid addition paths, offering more flexible positioning to meet complex requirements.

- ★ Fully enclosed robotic arm design, transmission parts have zero contact with acid vapors/liquids; no transmission lead screw required, consumable-free design, ensuring ultra-long service life.

- Tilted liquid addition tube design, reagents flow down the digestion tube wall, effectively reducing powder sample adhesion and preventing reagent/sample splashing.

- ★ Reagent usage prediction function: the instrument automatically calculates reagent requirements based on the method; if remaining reagent is insufficient, no liquid is added and software provides an early warning.

- After method completion, the instrument automatically flushes lines with water for automatic maintenance.

2.3 Fully Automatic Lifting & Oscillation System

- ★ Fully automatic high-frequency circular oscillation mixing function, using a mechanical motor oscillation system and digestion tube rack to mix reagents and samples, creating high-speed vortices within tubes to fully dissolve bottom and wall-adhered samples.

- Entire mixing process is contact-free with external environment, preventing contamination and cross-contamination; Can mix 84 digestion tubes simultaneously for higher efficiency.

- Oscillation power continuously adjustable, oscillation time settable; Multiple mixing modes available.

- ★ Dual-module design: The two modules can run completely independent digestion programs, enabling independent lifting/lowering, mixing, acid addition, different temperature digestion, acid driving, volume adjustment, etc.

- Automatic lifting function raises the tube rack before acid addition and after acid driving, lifting tubes clear of the heating module; Settable cooling time during lifted acid addition ensures safety.

2.4 Automatic Sample Volume Adjustment System

★ After digestion and acid driving, volume adjustment is performed at room temperature using short-range, low-power ultrasonic sensors to measure volume in the tubes; High-precision peristaltic and syringe pumps adjust samples to the specified volume (10-50mL); Adjustment accuracy: better than 0.5%.

2.5 Automatic Acid Fume Exhaust System

- Integrated automatic acid fume exhaust system, can be used without needing a fume hood. Contains fumes within the system, preventing damage to electronic components when placed entirely inside a hood.

- ★ Extra-large integrated fan, airflow up to 360 m³/h, ensures timely fume extraction and prevents condensation.

- ★ Exhaust system uses pneumatic-assisted opening, easily operated single-handedly; front cover self-locks in any position when open, preventing slippage.

- ★ Optimized small window design allows viewing experiments while preventing fume escape.

- ★ Standard exhaust system includes HEPA-grade filtration, continuously purifying air entering the system during digestion.

2.6 Intelligent Control Software

- PC-controlled via WiFi wireless connection, powerful functionality; Software supports easy switching between Chinese/English; Main interface shows reagent levels for 10 channels with alarms, and waste bottle full warning, with user-definable limits.

- ★ Scheduled start function: Methods can be set to run at a predefined time, the instrument starts digestion automatically, improving efficiency.

- ★ Voice prompt function: Notifies user upon method completion, preventing samples from drying due to residual heat, enhancing safety and reliability.

- ★ Remote monitoring system: Users can monitor real-time status and experimental steps from outside the lab (e.g., office) via phone or other terminals, tracking progress remotely; Allows remote emergency stop.

- ★ Smart maintenance reminders: Software automatically tracks instrument runtime, periodically prompting maintenance for filters, peristaltic pumps, etc.

- Freely set and store digestion and other pretreatment steps to meet various project requirements.

- Automatically generates digestion process reports.

3. Configuration List:

- One fully automatic graphite digestion main unit, including two independent graphite heating modules, 10-channel liquid addition/volume adjustment system, dual-module lifting/mixing module;

- One set of automatic acid fume exhaust system with HEPA purification;

- 5m polypropylene exhaust duct;

- One set of HEPA-grade filter purification mesh;

- Digestion tube racks (42-position × 2);

- One set of PTFE digestion tubes (84 tubes);

- One set of polypropylene digestion tubes (500 tubes);

- One set of intelligent control software;

- One brand laptop (Win 10 OS, 4G RAM, 500G HDD, 14");

- One set of remote monitoring system software; Optional

4. Technical Service & Instrument Training

- ★ The bid product must be mature, with over 400 domestic users; must provide 10 user acceptance reports (stamped) and a list of 400+ users.

- Engineers will install and debug the instrument at the user's lab upon notification until acceptance is passed.

- On-site free training for 2-5 technical personnel until independent operation is achieved, covering principles, debugging, operation, and maintenance.

- Manufacturer maintains an application lab domestically, providing free lifelong application technical consultation.

- ★ Manufacturer is ISO9001 certified, has an after-sales service center in ______, and responds within 24 hours to user repair requests.